**BIOFIL 3-month shower head filter**

Sterile

Reference: **20361**

Specification description

DELABIE anti-Legionella and all-germ shower head filter.

Device made from ABS.

Membrane type: tubular microfiltration.

Membrane made from polyethylene.

Pore cut-off threshold: 0.1μm nominal and 0.2μm absolute sterilising grade.

Qualified water volume per duration of use: 9,000 litres.

Filtration flow rate: 6.2 lpm at 1 bar, 11.2 lpm at 3 bar and 13.2 lpm at 5 bar (filter only, excludes any flow rate restriction in the tap).

Maximum upstream pressure at the point-of-use: 5 bar.

Maximum lifespan: 3 months after installation.

Withstands thermal shocks: 70°C for 30 minutes.

Withstands chemical shocks: 100ppm of active chlorine for a cumulative total of 3 hours.

Sterile filter.

CE Class Is medical device.

Weight: Empty filter 133g; Full filter 259g.

Packaged individually. Must be ordered in quantities of 10.

Packaged in boxes of 10 units.